# Automated Extraction of Nucleic Acid from Oral Fluid, Environmental Swabs and Fecal Samples with the KingFisher™ Flex

# Materials provided

| Contents | Cat. No. TC-9119-480 (480 reactions) | Storage |
| --- | --- | --- |
| VetAlert™ MagBead Proteinase K | 25 mL | 15-30°C (room temperature) |
| VetAlert™ MagBead Lysis Solution* | 190 mL |  |
| VetAlert™ MagBead Wash Buffer | 485 mL |  |
| VetAlert™ MagBead Elution Buffer | 45 mL |  |
| VetAlert™ MagBead Resin | 22 mL |  |
| VetAlert™ MagBead Stabilization Buffer | 100 mL |  |
| VetAlert™ MagBead Stabilization Additive | 2 mL | 2-8°C (refrigerated) |

*Note: If you see turbidity in the lysis solution, incubate the buffer at 37°C for 2 h in an incubator or water bath, and occasionally swirl the bottle until completely clear. Allow the buffer to return to room temperature before starting your protocol.

# Materials required but not provided

# Reagents

- 100% ethanol, ACS reagent grade or equivalent

- 100% isopropanol, ACS reagent grade or equivalent

- Molecular biology grade water

- (Optional) VetAlert™ DNA Internal Control System [Tetracore, Inc. Cat.# TC-9123-100, includes VetAlert™ DNA Internal Control (DNA IC) and VetAlert™ Internal Control Primer Probe Mix]

- (Optional) VetAlert™ RNA Internal Control System [Tetracore, Inc. Cat.# TC-9122-100, includes VetAlert™ RNA Internal Control (RNA IC) and VetAlert™ Internal Control Primer Probe Mix]

# Instrument and Equipment

- KingFisher™ Flex Purification System (Thermo Fisher Cat.# 5400630)

- Eppendorf ThermoMixer® C (Eppendorf Cat 5382000023) or similar and compatible microcentrifuge tubes.

- Eppendorf® SmartBlock Thermoblock for ThermoMixer® C & ThermoStat C (Eppendorf Cat. #5363000039) or similar and compatible Deepwell plate.

- Micropipettes and sterile pipette tips with aerosol barriers

- Laboratory mixer, Vortex or equivalent

- Microcentrifuge (Fisherbrand™ accuSpin™ Micro 17/17R Cat.#13-100-675) or similar and compatible Eppendorf tubes.

# TUBES, plates and other consumables

- KingFisher™ 96 deep-well plates (Thermo Fisher Cat.# 95040460)

- KingFisher™ 96 well microplate, 200 µL (Thermo Fisher Cat.# 97002540)

- KingFisher™ 96-tip comb for deep-well magnets (Thermo Fisher, Cat. #A43074)

- Mictrocentrifuge Tubes:1.5mL (Fisherbrand™ Cat.# 05-408-129)

- Adhesive PCR Plate Foils, or equivalent (Thermo Scientific™ Cat.# AB-0626)

# Before you begin

# Prepare Stabilization Solution:

- Prepare Stabilization Solution by combining 122.5 µL of Stabilization Buffer and 2.5 µL of Stabilization Additive per sample.

Note:  Prepare additional Stabilization Solution volume to accommodate pipetting loss.

# Prepare 50% Ethanol:

- Dilute 100% ethanol to 50% with an equal volume of molecular biology grade water.

# Preprocessing of samples in microcentrifuge tubes

- Transfer 250 µL of sample to a microcentrifuge tube.

- Add 125 µL of the prepared Stabilization Solution to each sample.

- Add 350 µL of Lysis Solution to each sample.

Optional: Add 6 µL of VetAlert™ RNA IC or VetAlert™ DNA IC to the Lysis Solution if IC is used as an extraction control.

- Mix sample on an Eppendorf Thermomixer (or similar) at 1400 rpm for 3 minutes.

Centrifuge sample tubes at approximately 16,000 x g for 2 minutes.

# Preprocessing of samples in a deep-well plate

- Transfer 250 µL of sample to the appropriate wells of a deep-well plate.

- Add 125 µL of the prepared Stabilization Solution to each sample well.

- Add 350 µL of Lysis Solution to each sample well.

Optional: Add 6 µL of VetAlert™ RNA IC or VetAlert™ DNA IC to the Lysis Solution if IC is used as an extraction control.

- Seal the deep-well plate with sealing foil and shake the plate samples on an Eppendorf Thermomixer (or similar) at 1400 rpm for 5 minutes.

- Centrifuge sample in deep-well plate at ≥ 3000 x g for 5 minutes.

# Prepare purification plate

- Wash Plate 1: Add 100 µL of 50% ethanol and 900 µL of Wash Buffer to each well of a KingFisher™ 96 deep-well plate required for purification.

- Wash Plate 2: Add 450 µL of 50% ethanol to each well of a KingFisher™ 96 deep-well plate required for purification.

- Elution Plate: Add 50 µL of Elution Buffer to each well of a KingFisher™ 96 well microplate (200 µL) required for purification.

- Tip Comb Plate: Place KingFisher™ 96 tip comb into an empty KingFisher™ 96 well microplate (200 µL).

- Sample plate

- Add 30 µL of Proteinase K to a new KingFisher™ 96 deep-well plate, in accordance with your sample processing plate map.

- Gently transfer 600 µL of sample lysate supernatant from your sample processed plate containing sample lysate into the new deep-well that contains Proteinase K plate.

- Add 350 µL of 100% isopropanol to each well containing sample.

- Add 20 µL of Resin to each well containing sample.

Note: The magnetic beads tend to drop out of suspension when not in use. Please ensure the Resin is completely resuspended prior to transfer by vigorously shaking or vortexing the bottle.

# Begin process on the KingFisher Flex

- Start KingFisher method Tetracore_VetAlert_MagBeadExtraction.bdz and follow instrument prompts for plate placement (see below).

| Plate ID | Plate position on the KingFisher Flex |
| --- | --- |
| Sample Plate | 1 |
| Wash Plate 1 | 2 |
| Wash Plate 2 | 3 |
| Elution Plate | 4 |
| Tip Comb Plate | 5 |

- Follow prompts for removal of plates from the KingFisher Flex.

- Transfer final eluates to a container of your choice for storage.

Note: Extracted nucleic acid can be stored at 4°C (2°C to 8°C) for immediate use. For long-term storage DNA should be stored at -20°C (-15°C to -25°C) and RNA at -80°C (-60°C to -90°C).

# Symbols

| Symbol | Meaning | Symbol | Meaning |
| --- | --- | --- | --- |
|  | Catalog Number |  | Lot Number |
|  | Consult Instructions for Use |  | Manufacturer |
|  | Expiration Date |  | Temperature Limit |

# Revision History

| Revision | Date | Description of Change |
| --- | --- | --- |
| 00 | 18 July 2024 | Initial Release |
| 01 | 2 February 2026 | Revision |

Manufactured by: 	Tetracore, Inc.

77 Upper Rock Circle

Suite 600

Rockville, MD 20850

U.S.A.

Phone:  240.268.5400

www.tetracore.com

Tetracore® is a registered trademark of Tetracore, Inc.

VetAlert™ is a Tetracore trademark.

Revision: 02 February 2026

Part number:  PLM-0799-1